Robust computational spectral sensing via orthogonal frequency encoding

Yijia Zeng1, Zhi Zhang1, Quanzhou Long1, Yafan Jin1, Zijian Lin1, Feng Zhou 1, Tingbiao Guo1,2,*, Sailing He1,2,*

1 Centre for Optical and Electromagnetic Research, National Engineering Research Center for Optical Instruments, College of Optical Science and Engineering, Zhejiang University, Hangzhou, Zhejiang, China

2 Institute of Fundamental and Transdisciplinary Research, Zhejiang University, Hangzhou, Zhejiang, China

3 Department of Electromagnetic Engineering, School of Electrical Engineering, Royal Institute of Technology, Stockholm S-10044, Sweden

*Corresponding authors: sailing@kth.se and tbguo@zju.edu.cn

Abstract

High-fidelity spectroscopy in real-world environments is severely limited by unpredictable ambient interference and the inherent trade-off between acquisition speed and signal-to-noise ratio (SNR). Conventional time-division multiplexing approaches often fail as they lack intrinsic mechanisms to decouple target signals from background illumination. Here, we propose orthogonal frequency encoding, a robust framework that shifts spectral acquisition from serial scanning to a parallelized frequency-domain architecture. By encoding each spectral channel with a distinct orthogonal frequency kernel, it allocates carriers to frequency gaps to evade structured environmental noise. As a proof of concept, we use a programmable LED ·array to demonstrate single-shot spectral reconstruction based on this framework. The system exhibits immunity to high-energy interference, maintaining high reconstruction fidelity (RMSE ≈ 0.0261) under an SNR of -11.77 dB. This work establishes a fast, low-cost strategy for spectral sensing in challenging environments.

Introduction

Spectroscopic sensing stands as an essential technology across diverse disciplines[1], finding critical utility in precision agriculture[2], environmental monitoring[3], and clinical diagnostics[4]. However, achieving high-fidelity spectroscopy in real-world environments remains a significant challenge due to unpredictable ambient interference. Conventional dispersive spectrometers based on gratings or prisms are not only compromised by bulky form factors and high costs but also inherently constrained by the need for narrow entrance slits, which limits optical throughput[5]. Moreover, these systems indiscriminately accumulate total photon flux due to a lack of intrinsic mechanisms to decouple the target signal from background illumination.[6] This mixes the target signal with ambient interference, causing spectral reconstruction to fail in real-world scenarios.

Computational spectroscopy offers a promising pathway to overcome these physical limitations[7]. Existing schemes remain constrained by an inherent trade-off between signal-to-noise ratio (SNR)[8], acquisition speed[9], and spectral resolution[10]. Snapshot approaches[11], such as filter arrays, achieve high acquisition speed but suffer from a spatial-spectral trade-off that limits their spectral resolution[12]. Conversely, TDM methods rely on sequential scanning, trading off temporal resolution for spectral accuracy. This sequentiality dictates that the total acquisition time scales linearly with the number of spectral channels (Ttotal∝N), rendering them ineffective for capturing high-speed dynamics[13]. Moreover, these computational frameworks remain bound to controlled lighting environments due to their lack of an intrinsic encoding mechanism to effectively decouple the active signal from ambient interference.

In this work, we propose a robust framework for computational spectral sensing in challenging environments, termed orthogonal frequency encoding (OFE). Unlike standard lock-in detection or post-detection unmixing strategies[14], this scheme ensures isolation by encoding each spectral basis function via a distinct orthogonal frequency kernel, enabling the simultaneous reconstruction of all encoded spectral channels via a detector in a single-shot acquisition. We employ a low-cost, programmable LED array as a demonstration of this framework.	The results demonstrate that this strategy enables high-fidelity spectral reconstruction at an SNR of -11.77 dB. Furthermore, it achieves fast sensing with acquisition times as short as 250 μs. This work establishes a robust, fast and low-cost framework for spectral sensing in challenging environments.

Principle

A generalized spectral sensing architecture can be abstracted into two fundamental components: a spectral modulator and a detector. Under a conventional time-division multiplexing (TDM) scheme, spectral channels are activated sequentially according to a predefined timing sequence. The instantaneous intensity yTDMt received by the detector at any given time t represents the superposition of the spectral signal from the single channel within the current time slot and the ambient noise:

|  | yTDMt=k=1Kλ1λ2PkλRλSλdλ⋅rectt-kττ+Inoiset | (1) |
| --- | --- | --- |

where Pkλ denotes the spectral basis function of the k-th channel for the spectral sensing system. rect⋅ represents the unit rectangular function, acting as a temporal gating operator that characterizes the serial scanning mechanism, where the k-th channel is active solely during its specific integration time τ. Rλ represents the reflectance or transmittance spectrum of the target, Sλ denotes the detector’s spectral responsivity, encompassing both photoelectric conversion efficiency and transimpedance gain, Inoiset denotes the ambient light noise.

In the orthogonal frequency encoding scheme, we substitute the binary temporal gating operator rect⋅ with a modulation kernel mkt;fk uniquely assigned to the k-th channel. This substitution eliminates the requirement for sequential time slots, allowing all channels to operate in parallel. The modulation kernel for the k-th channel is formulated as:

|  | mkt;fk=ak,0+ak,1cos2πfkt+φk | (2) |
| --- | --- | --- |

where fk is the distinct carrier frequency assigned to channel k. ak,0 represents the direct current (DC) bias required to ensure non-negative optical intensity. ak,1 denotes the carrier amplitude, and φk is the initial phase of the signal.

Upon simultaneous collection of all spectral channels modulated by their respective kernels, the detector integrates the total photon flux. The received OFE signal yOFEt is a composite waveform containing information from all K channels concurrently:

|  | yOFEt=k=1Kλ1λ2PkλRλSλdλ zk⋅mkt;fk+Inoiset | (3) |
| --- | --- | --- |

To simplify the physical interpretation, we define the scalar zk as the spectral projection value for the k-th channel. Substituting the modulation kernel into equation (3), the received signal yOFEt can be reorganized into two distinct components:

|  | yOFEt=k=1Kzkak,0+InoisetDC & Noise component+k=1Kzkak,1cos2πfkt+φkSignal component | (4) |
| --- | --- | --- |

The signal decomposition is achieved by mapping the time-domain observation into the frequency domain. Defining F{⋅} as the continuous Fourier transform operator, the frequency-domain response Yf is given by:

|  | Yf=F{Inoiset}+k=1K12zkak,1ejφkδf-fk+e-jφkδf+fk | (5) |
| --- | --- | --- |

where δ⋅ denotes the Dirac delta function. This formulation demonstrates the mechanism of noise rejection: if the modulation frequencies {fk} are designed to avoid the dominant bands of environmental noise without mutual interference, the projection value zk is retrieved by extracting the magnitudes at the corresponding carrier frequencies:

|  | Yfk=12zkak,1⇒zk=2Yfkak,1 | (6) |
| --- | --- | --- |

To ensure strict orthogonality and avoid mutual interference within a physical optoelectronic system, the carrier allocation must satisfy three boundary conditions: spectral separability (fi-fj>Δf, where  Δf is the frequency resolution), harmonic interference avoidance (fj≠n⋅fi), and intermodulation avoidance (fi±fj≠fk). Mathematical derivations of these constraints and the corresponding switch search algorithm are detailed in Supplement Sections 2 and 3.

Figure 1. OFE schematic diagram. The signal chain initiates with the modulation of discrete spectral basis functions using temporally orthogonal sinusoidal kernels. These components are superimposed into a single time-domain composite signal, which is subjected to environmental noise during physical transmission. Following sensor reception, a mathematical transform (e.g., Fourier transform) maps the noise-contaminated temporal waveform into the frequency domain. This transformation isolates the deterministic carrier peaks from the stochastic noise floor, enabling a computational algorithm to reconstruct the target spectra from these high-fidelity frequency-domain projections.

OFE possesses two advantages over traditional TDM: the multiplex advantage and amplitude immunity advantage. The multiplex advantage stems from its parallel collection over K channels. Coherent frequency-domain demodulation acts as a matched filter, bounding the noise variance to σOFE2∝1/K. This inherent variance ratio of K mathematically dictates a reduction in demodulation error by a factor of K under identical acquisition time. This theoretical multiplexing gain, conceptually analogous to Fellgett’s advantage, preserves high-fidelity reconstruction in additive white Gaussian noise (AWGN) environments. Amplitude immunity advantage lies in encoding information in the frequency rather than amplitude domain, facilitating the physical excision of amplitude-based noise components.

Once the projection values for all K channels are retrieved, they are aggregated to form the measurement vector z=z1,z2,…,zKT. To recover the spectrum Rλ of the target, we discretize the spectral range λ1,λ2 into N bands and model the measurement process as a system of linear equations:

|  | z=Wr+e | (7) |
| --- | --- | --- |

where r∈RN represents the ground truth (GT) spectrum, and W∈RK×N is the pre-calibrated system matrix determined by the spectral sensing system. The term e accounts for the system noise and measurement residuals. Consequently, spectral reconstruction is inherently an ill-posed inverse problem. To ensure physical fidelity and numerical stability, we recast the reconstruction as a constrained convex optimization problem. We employ a projected gradient descent (PGD) algorithm to minimize a composite objective function that incorporates both physical boundary constraints and adaptive spectral priors:

|  | r=min0⪯r⪯112z-Wr22+Rr | (8) |
| --- | --- | --- |

where Rr is a generalized regularization term adaptively selected based on the physical characteristics of the target spectra. For targets with distinct narrowband features, we apply a sparsity prior to resolve fine spectral structures, defining Rr=β1r. For targets characterized by continuous, broadband spectral envelopes, we enforce smoothness and structural continuity by defining Rr=β22r22+β32Dr22, where D∈RN-1×N is the first-order difference operator. Detailed hyperparameter configurations and the specific algorithmic implementation are cataloged in Supplement Section 1.

Experiment

To isolate and quantify the fundamental multiplexing advantage of the OFE framework against AWGN, we establish a simulation benchmark. Operating across the visible spectrum (400–700 nm), the architecture deploys K=12 discrete spectral channels. Rather than employing idealized Gaussian approximations, the spectral basis functions are directly instantiated from the empirically calibrated radiometric profiles of our physical LED array (detailed in Supplement Section 6, Fig. S2). This ensures that the system matrix W encapsulates the non-ideal spectral irregularities and native radiometric variations inherent to the physical hardware. In the time domain, these channels are modulated by the orthogonal composite sinusoidal carriers. The numerical simulations are conducted utilizing the CAVE multispectral database[15].

In a pure AWGN regime with a flat noise power spectral density, the specific spectral placement of orthogonal carriers mathematically yields identical variance reduction. Therefore, to ensure baseline orthogonality without introducing redundant hardware-specific combinatorial complexities, we employ a straightforward sub-octave linearly spaced carrier set for all numerical TDM-vs-OFE comparative simulations. Specifically, this frequency set (Strategy 1) spans from fmin=380 kHz to fmax=490 kHz with a uniform frequency internal of 10 kHz (we term this frequency selection as strategy 1, and detailed algorithmic implementation is provided in Supplement Section 3). By the condition of orthogonal integration (T = 1/Δf), this 10 kHz spectral resolution dictates a 100 μs acquisition window. To ensure an equivalent temporal and energy budget for comparison, the TDM benchmark uniformly partitions the identical 100 μs acquisition window into K=12 sequential time slots (≈8.33 μs per channel), with each channel operating at a constant direct current (DC) baseline. To mitigate computational redundancy while preserving statistical robustness, the errors reported in Fig. 2 are evaluated over a 10% globally random subset of pixels from all 29 images in the CAVE dataset.

Figure 2. Numerical benchmarking of the OFE and TDM frameworks. (a) Time-domain waveforms with AWGN (SNR = 25 dB). (b) Spectral reconstruction performance: RMSE (top) and demodulation error (bottom) versus SNR. Data points represent statistical averages across the CAVE database. (c) Time-domain waveforms under multitone interference (Δfguard=10 kHz, SNR = -10 dB). The 10 kHz offset corresponds to the 1/T frequency resolution limit of the sampling window. (d) RMSE versus multitone interference SNR at varied frequency offsets. TDM yields a single baseline (grey dashed line) due to the absence of frequency-domain selectivity. OFE achieves interference rejection exclusively when Δfguard meets the 10 kHz integer-cycle threshold.

Figure 2(a) illustrates the time-domain waveforms of both schemes under AWGN (SNR = 25 dB) within a 20 μs window. Note that a DC bias (ak,0) is added to the OFE waveform, presenting the zero-mean information-bearing composite signal. The OFE signal exhibits a continuous, dense composite waveform characteristic of parallel frequency multiplexing, whereas TDM presents discrete, sequential step-like features. Figure 2(b) quantifies the reconstruction accuracy across a broad SNR spectrum. To directly link the hardware acquisition quality to the algorithmic reconstruction, we track both the spectral reconstruction RMSE and the demodulation error defined as the relative Euclidean distance z-z2/z2 between the extracted and ground-truth (GT) spectral projections). Across all evaluated noise levels, OFE achieves lower errors than TDM. The observed separation between the curves aligns with the theoretical K (approximately 3.46-fold for K=12) multiplexing advantage derived in our principle section.

Beyond broadband AWGN, practical systems frequently encounter structured interference. To isolate the system’s dynamic response to such artifacts, we maintain the identical TDM and OFE configurations used in the AWGN benchmark, but substitute the broadband noise model with severe multitone interference operating at an aggressive SNR of -10 dB. We parameterize the frequency offset between each interfering tone and its adjacent OFE carrier as the guard band, Δfguard=fnoise-fcarrier.

Figures 2(c) and 2(d) elucidate the interference rejection mechanism. In the time-domain visualization (Fig. 2c), multitone interference (SNR = -10 dB) at an offset of Δfguard=10 kHz induces high-frequency ripples. Because this offset matches the frequency resolution limit (1/T) of the acquisition window, the noise harmonic completes integer cycles within the integration period. This satisfies the orthogonality condition, enabling its complete elimination during frequency-domain demodulation. Figure 2(d) quantifies this threshold behavior. TDM lacks frequency-domain selectivity; consequently, its time-domain integration degrades uniformly across all evaluated offsets, collapsing into a single baseline (grey dashed line). For OFE, noise immunity is governed by the 1/T (1/100 μs) threshold. Offsets below this limit (Δfguard=1 and 5 kHz) fail to satisfy the coherent integration constraint, resulting in non-zero residual integrals that manifest as spectral leakage and elevated RMSE. Once the offset reaches the 10 kHz threshold, orthogonality is restored, yielding a consistently low reconstruction error independent of the interference intensity.

Following the simulation-based verification of the theoretical noise immunity, we develop a prototype to validate the OFE framework experimentally. To hardware-instantiate the K=12 spectral channels modeled in the numerical benchmark, we deploy the aforementioned programmable LED array as the physical spectral encoder. Spanning the visible band (400–700 nm), this illumination source provides the discrete radiometric profiles (detailed in Supplement Section 6, Fig. S2) necessary for orthogonal modulation. Following spectral interaction with the target, a photodiode coupled to a high-speed oscilloscope captures the time-domain composite waveform. Digital demodulation is subsequently executed via the fast Fourier transform (FFT), mapping the temporal signal into the frequency domain to extract the orthogonal spectral projections. This configuration ensures that the physical system matrix W maps to our theoretical benchmark. Figure 3(a) schematizes the architectural configurations for both reflective and transmissive modalities. Comprehensive hardware specifications, including the physical bandwidth constraints, alongside the benchtop photograph are detailed in Supplement Section 7, Fig. S3.

To assess the baseline reconstruction capabilities without external electrical interference, we configure the hardware for a high-speed acquisition (Ts=250 μs). As established by the empirical benchmark (Fig. S4 in Supplement Section 9), this temporal boundary dictates a theoretical frequency resolution of Δf = 4 kHz. To multiplex all K=12 carriers within the available hardware bandwidth under this constraint, we deploy the broadband greedy search algorithm (Strategy 2). Unlike the dense, linearly spaced configuration used in the pure AWGN benchmark, this strategy yields a non-uniformly distributed carrier set (Hopt∈30,410 kHz) that structurally navigates the complex intermodulation landscape, maximizing spectrum utilization in a clean electromagnetic environment (the exact numerical array is detailed in Supplement Section 3).

Figures 3(b) and 3(c) show the spectral reconstruction fidelity under reflective (Fig. 3b) and transmissive (Fig. 3c) modes. Figure 3(b) demonstrates the recovery of continuous broadband envelopes from a standard 24-patch reflective colorchecker. The system yields a mean RMSE of 0.0422, which translates into consistent sRGB colorimetric reproduction (inset). Figure 3(c) validates the recovery of isolated 10 nm FWHM narrowband filters. The extracted spectral vectors successfully resolve these sharp peaks, preserving the physically discontinuous features. While the large discrepancy at some peaks may be due to the limited number of spectra basis functions (K=12) whose overlapping spectral responses render the inverse problem locally ill-conditioned, biasing reconstructed energy toward spectrally denser regions.

Figure 3. Experimental validation of the OFE prototype in the high-speed regime (Ts=250 μs, Strategy 2). (a) Schematics of the optical paths for (i) reflective and (ii) transmissive modalities. (b) Reconstructed surface reflectance of the 24-patch reflective colorchecker (odd-numbered patches from #1 to #23, indexed sequentially left-to-right, top-to-bottom). Open circles denote the OFE prediction (Pred.) against the ground truth (GT, dashed lines), yielding an average RMSE of 0.0422. Insets display the corresponding sRGB color renderings. (c) Reconstruction profiles of isolated 10 nm FWHM narrowband spectral peaks.

Furthermore, we introduce a 50 kHz switched-mode power supply (SMPS) noise source to evaluate its amplitude immunity advantage with the setup shown in Figure 3a(ii). Patch #15 in a transmissive colorchecker is selected as a representative broadband profile. As depicted in the frequency-domain analysis (Figure 4d), the SMPS injects deterministic harmonics that dominate the low-frequency spectrum up to approximately 350 kHz. Maintaining the broadband configuration (Strategy 2, 30–410 kHz) from Figure 3 would result in spectral collisions for the lower-frequency carriers. To actively evade this noise, we dynamically reconfigure the hardware to the sub-octave configuration (Strategy 1), explicitly shifting all OFE carriers, spaced at 10 kHz intervals, into the higher 380–490 kHz band (Fig. 4c). Consequently, as dictated by the empirical benchmark (Fig. S4 in Supplement Section 9), maintaining orthogonality for this dense frequency cluster necessitates prolonging the temporal integration window to Ts=1000 μs.

Figure 4. Experimental robustness evaluation under 50 kHz SMPS interference (Ts=1000 μs, Strategy 1). Time-domain composite signals under (a) high SNR (23.11 dB) and (b) extreme low SNR (-11.77 dB). (c) Frequency-domain spectrum of the clean OFE signal, with red triangles marking the sub-octave carriers (380–490 kHz). (d) Frequency-domain spectrum of the interference environment, illustrating the physical isolation between the low-frequency SMPS harmonics and the high-frequency OFE signal area. (e) Reconstruction RMSE versus SNR, demonstrating a robustness plateau up to the -12 dB threshold. (f) Reconstructed transmittance spectra across varying noise levels against the ground truth (solid black line).

To quantify the system’s performance, we define an effective frequency-domain SNR. As detailed in Supplement Section 8, this metric compares the power of the fundamental carriers against the broad environmental noise floor. By excluding the deterministic, higher-order hardware harmonics from the noise calculation, we isolate the true stochastic impact of the SMPS interference, providing an assessment of the framework’s environmental robustness.

Figures 4(a) and 4(b) contrast the time-domain composite signals under pristine (SNR = 23.11 dB) and extreme interference (SNR = -11.77 dB) conditions within a 100 μs window. In Figure 4(b), the high-energy SMPS noise entirely masks the periodic modulation envelope. However, the system’s resilience is grounded in the strict frequency-domain separation depicted in Figures 4(c) and 4(d). Because we dynamically deployed the sub-octave configuration (Strategy 1), the OFE carriers remain rigorously isolated within the higher-frequency 380–490 kHz band (Fig. 4c). This structural frequency isolation ensures that despite the severe time-domain distortion observed in Fig. 4b, the orthogonal carriers remain uncontaminated prior to the demodulation stage.

Figure 4(e) quantifies the aggregate statistical tolerance to this severe environmental noise. The mean RMSE exhibits a distinct robustness plateau (0.0261) from 23.11 dB to even -11.77 dB, with performance breakdown initiating only when the SNR breaches around the -12 dB threshold. To visualize the underlying structural degradation trajectory at this boundary, Figure 4(f) maps the reconstructed spectra across the evaluated noise levels. Down to the extreme -11.77 dB threshold, the extracted curves strictly track the ground truth (solid black line), validating the framework’s robustness within the sub-noise regime.

Conclusion and discussion

In this work, we propose OFE architecture to resolve the speed-SNR trade-off inherent in conventional TDM. By shifting from serial scanning to parallelized frequency-domain encoding, the system achieves single-shot spectral reconstruction in 250 μs. Crucially, unlike standard lock-in techniques[16], OFE employs a strategic frequency allocation scheme that exploits the spectral gaps of structured environmental noise. This mechanism ensures robust retrieval (RMSE ≈ 0.0261) even under harsh interference with an SNR of -11.77 dB (at a 1000 μs sampling time), effectively immunizing the system against common artifacts like PWM dimming and SMPS ripples.

However, the performance boundaries of OFE are strictly dictated by hardware constraints rather than the encoding strategy itself. The acquisition speed T is governed by the time-frequency uncertainty principle, T ≈1/Δf. While our prototype is limited to 250 μs by the 500 kHz bandwidth of LED drivers, the framework holds immense potential for ultrafast spectroscopy. By transitioning to high-speed laser diodes modulated at hundreds of megahertz or even gigahertz, the system can theoretically scale to nanosecond-level acquisition times[17], allowing for a dramatic increase in the number of spectral channels without compromising temporal resolution.

To extend the system’s capabilities, the OFE architecture offers versatile scalability across spectral, spatial, and functional dimensions. First, the spectral range is not confined to the visible band; by incorporating appropriate sources, modulators and detectors, the framework can be readily expanded to the ultraviolet or infrared regions[18], enabling molecular fingerprinting in complex environments. Second, coupling OFE with event cameras[19] (capable of microsecond-level temporal resolution) or SPAD arrays[20] (achieving picosecond to nanosecond-level precision) facilitates hyperspectral imaging. Finally, the active nature of OFE opens new possibilities for real-time adaptive sensing. Unlike passive spectrometers, the spectral channels in OFE can be dynamically optimized based on initial measurements to maximize the reconstruction fidelity for specific spectral features[21], establishing a closed-loop framework for intelligent optical sensing.

Acknowledgement

We acknowledge the support from the Ningbo Public Welfare Research Program Project (2024Z234), the “Pioneer and Leading Goose” R&D Program of Zhejiang (Nos. 2025C02159, 2025C02140, 2025C02159, 2023C03083, 2024C03045, 2024Z146), the National Natural Science Foundation of China (W2412107, 91233208, 62105284), the National Key Research and Development Program of China (2022YFB2804100, 2022YFC3601002), Ningbo Science and Technology Project (Nos. 2020Z077, 2023Z179, 2023Z122, 2024Z146), Science and Technology Plan Key Project of Taizhou City (24gyz01), Ningbo Top Medical and Health Research Program (No.2022020506), the Special Development Fund of Shanghai Zhangjiang Science City, the Special Development Fund of Hangzhou Chengxi Sci-tech Innovation Corridor, and the Fundamental Research Funds for the Central Universities (226-2024-00150). National Ten Thousand Talent Program (Young Talents)

References

- Cui X, Nigmatulin F, Wang L, et al. Miniaturized spectral sensing with a tunable optoelectronic interface[J]. Science advances, 2025, 11(4): eado6886.

- Zhang D, Hou L, Lv L, et al. Precision agriculture: temporal and spatial modeling of wheat canopy spectral characteristics[J]. Agriculture, 2025, 15(3): 326.

- Oberski T, Walendzik B, Szejnfeld M. The monitoring of macroplastic waste in selected environment with UAV and multispectral imaging[J]. Sustainability, 2025, 17(5): 1997.

- Zou X, Shen Y, Jiao C, et al. Label-free resonance Raman spectral imaging reveals magnesium microsphere therapy attenuates oxidative damage in knee osteoarthritic rats[J]. Spectrochimica Acta Part A: Molecular and Biomolecular Spectroscopy, 2026: 127743.

- Liu W, Xu S, Lee C. Ultracompact on-chip spectral shaping using pixelated nano-opto-electro-mechanical gratings[J]. Science, 2025, 389(6762): 806-810.

- Grotevent M J, Yakunin S, Bachmann D, et al. Integrated photodetectors for compact Fourier-transform waveguide spectrometers[J]. Nature Photonics, 2023, 17(1): 59-64.

- Hong D, Li C, Yokoya N, et al. Hyperspectral imaging[J]. Nature Reviews Methods Primers, 2026, 6(1): 19.

- Cui K, Huang Y. On-chip spectral imaging and sensing transition towards marketable technologies[J]. Nature Reviews Electrical Engineering, 2025, 2(3): 151-152.

- Chen Z, Qu G, He Z, et al. Single‐Pixel Multispectral Imaging With Customized Pixelated Spectral Mask Plate[J]. Laser & Photonics Reviews, 2026: e03207.

- Hong D, Zhang B, Li X, et al. SpectralGPT: Spectral remote sensing foundation model[J]. IEEE transactions on pattern analysis and machine intelligence, 2024, 46(8): 5227-5244.

- Hua X, Wang Y, Wang S, et al. Ultra-compact snapshot spectral light-field imaging[J]. Nature communications, 2022, 13(1): 2732.

- Yao Z, Liu S, Wang Y, et al. Integrated lithium niobate photonics for sub-ångström snapshot spectroscopy[J]. Nature, 2025, 646(8085): 567-575.

- Zhou P, Lv Y, Zheng Y, et al. CTIS spectral image reconstruction technology based on slit-scanning architecture[J]. Optics Express, 2024, 32(14): 25276-25291.

- Adhikari S, Gross N, Wilson K S, et al. On-chip lock-in detection for ultrafast spectroscopy of single particles[J]. The Journal of Physical Chemistry C, 2024, 128(21): 8708-8715.

- J. I. Park, M. H. Lee, M. D. Grossberg, et al. 2007 IEEE 11th International Conference on Computer Vision. IEEE, 2007: 1-8.

- Shao Z, Huang G, Mielczarek A, et al. Robust spectral sensor for standoff biometric detection[J]. Nature Sensors, 2026: 1-8.

- Wörner H J, Wolf J P. Ultrafast spectroscopy of liquids using extreme-ultraviolet to soft-X-ray pulses[J]. Nature Reviews Chemistry, 2025, 9(3): 185-199.

- Yu H, Memon M H, Yao M, et al. A miniaturized cascaded-diode-array spectral imager[J]. Nature Photonics, 2025, 19(12): 1322-1329.

- Yu B, Liang J, Wang Z, et al. Active hyperspectral imaging using an event camera[C]//Proceedings of the Computer Vision and Pattern Recognition Conference. 2025: 929-939.

- Cao Y, Bao H, Peng L, et al. Single-Photon-Camera-Based Time and Spatially Resolved Electroluminescence Spectroscopy for Micro-LED Analysis[J]. ACS Photonics, 2024, 11(2): 731-736.

- Yuan Y, Sun R, Xie X, et al. Multispectral imaging colorimetric measurement method with illumination-adaptive white calibration[J]. JOSA A, 2025, 42(7): 898-907.