Robust spectral sensing via orthogonal frequency encoding

Yijia Zeng1, Zhi Zhang1, Quanzhou Long1, Zijian Lin1, Feng Zhou 1, Tingbiao Guo1,2,*, Sailing He1,2,*

1 Centre for Optical and Electromagnetic Research, National Engineering Research Center for Optical Instruments, College of Optical Science and Engineering, Zhejiang University, Hangzhou, Zhejiang, China

2 Institute of Fundamental and Transdisciplinary Research, Zhejiang University, Hangzhou, Zhejiang, China

3 Department of Electromagnetic Engineering, School of Electrical Engineering, Royal Institute of Technology, Stockholm S-10044, Sweden

*Corresponding authors: sailing@kth.se and tbguo@zju.edu.cn

Abstract

High-fidelity spectroscopy in real-world environments is severely limited by unpredictable ambient interference and the inherent trade-off between acquisition speed and signal-to-noise ratio (SNR). Conventional time-division multiplexing approaches often fail as they lack intrinsic mechanisms to decouple target signals from background illumination. Here, we propose orthogonal frequency encoding, a robust framework that shifts spectral acquisition from serial scanning to a parallelized frequency-domain architecture. By encoding each spectral channel with a distinct orthogonal frequency kernel, it allocates carriers to frequency gaps to evade structured environmental noise. As a proof of concept, we use a programmable LED array to demonstrate single-shot spectral reconstruction based on this framework. The system exhibits immunity to high-energy interference, maintaining high reconstruction fidelity (RMSE ≈ 0.0261) under an SNR of

-11.77 dB. This work establishes a fast, low-cost strategy for spectral sensing in challenging environments.

Introduction

Spectroscopic sensing stands as an essential technology across diverse disciplines[1], finding critical utility in precision agriculture[2], environmental monitoring[3], and clinical diagnostics[4]. However, achieving high-fidelity spectroscopy in real-world environments remains a significant challenge due to unpredictable ambient interference. Conventional dispersive spectrometers based on gratings or prisms are not only compromised by bulky form factors and high costs but also inherently constrained by the need for narrow entrance slits, which limits optical throughput[5]. Moreover, these systems indiscriminately accumulate total photon flux due to a lack of intrinsic mechanisms to decouple the target signal from background illumination.[6] This mixes the target signal with ambient interference, causing spectral reconstruction to fail in real-world scenarios.

Computational spectroscopy offers a promising pathway to overcome these physical limitations[7]. Existing schemes remain constrained by an inherent trade-off between signal-to-noise ratio (SNR)[8], acquisition speed[9], and spectral resolution[10]. Snapshot approaches[11], such as filter arrays, achieve high acquisition speed but suffer from a spatial-spectral trade-off that limits their spectral resolution[12]. Conversely, TDM methods rely on sequential scanning, trading off temporal resolution for spectral accuracy. This sequentiality dictates that the total acquisition time scales linearly with the number of spectral channels (Ttotal∝N), rendering them ineffective for capturing high-speed dynamics[13]. These computational frameworks remain bound to controlled lighting environments due to their lack of an intrinsic encoding mechanism to effectively decouple the active signal from ambient interference.

In this work, we propose a robust framework for spectral sensing in challenging environments, termed orthogonal frequency encoding (OFE). Unlike standard lock-in detection or post-detection unmixing strategies[14], this scheme ensures isolation by encoding each spectral basis function via a distinct orthogonal frequency kernel, enabling the simultaneous reconstruction of all encoded spectral channels via a detector in a single-shot acquisition. We employ a low-cost, programmable LED array as a demonstration of this framework.	The results demonstrate that this strategy enables high- fidelity spectral reconstruction at an SNR of -11.77 dB. Furthermore, it achieves fast sensing with acquisition times as short as 250 μs. This work establishes a robust, fast and low-cost framework for spectral sensing in challenging environments.

Principle

Under conventional time-division multiplexing (TDM) scheme, spectral channels are activated sequentially according to a predefined timing sequence. The instantaneous intensity yTDMt received by the detector at any given time t represents the superposition of the spectral signal from the single channel within the current time slot and the ambient noise:

|  | yTDMt=k=1Krectt-kττλ1λ2Pkλk=1KPkλ⋅rectt-kττ⋅RλSλdλ+Inoiset | (1) |
| --- | --- | --- |

where Pkλ denotes the spectral basis function of the k-th channel for the spectral sensing system. rect⋅ represents the unit rectangular function, acting as a temporal gating operator that characterizes the serial scanning mechanism, where the k-th channel is active solely during its specific integration time τ. Rλ represents the reflectance or transmittance spectrum of the target, Sλ denotes the detector's spectral responsivity, encompassing both photoelectric conversion efficiency and transimpedance gain. And Inoiset denotes the ambient light noise.

Figure 1. OFE schematic diagram.

In orthogonal frequency encoding scheme, we substitute the binary temporal gating operator rect⋅ with a modulation kernel mkt;fk uniquely assigned to that channel. This substitution eliminates the requirement for sequential time slots, allowing all channels to operate in parallel. The modulation kernel for the k-th channel is formulated as:

|  | mkt;fk=ak,0+ak,1cos2πfkt+φk | (2) |
| --- | --- | --- |

where fk is the distinct carrier frequency assigned to channel k. ak,0 represents the direct current (DC) bias required to ensure non-negative optical intensity. ak,1 denotes the carrier amplitude, and φk is the initial phase of the signal.

Upon simultaneous projection of all spectral channels modulated by their respective kernels, the detector integrates the total photon flux. The received OFE signal yOFEt is a composite waveform containing information from all K channels concurrently:

|  | yOFEt=k=1Kλ1λ2PkλRλSλdλ zk⋅mkt;fk+Inoiset | (3) |
| --- | --- | --- |

To simplify the physical interpretation, we define the scalar zk as the spectral projection value for the k-th channel. Substituting the modulation kernel into equation (3), the received signal yOFEt can be reorganized into two distinct components:

|  | yOFEt=k=1Kzkak,0+InoisetDC & Noise component+k=1Kzkak,1cos2πfkt+φkSignal component | (4) |
| --- | --- | --- |

The signal decomposition is achieved by mapping the time-domain observation into the frequency domain. Defining F{⋅} as the continuous Fourier transform operator, the frequency-domain response Yf is given by:

|  | Yf=F{Inoiset}+k=1K12zkak,1ejφkδf-fk+e-jφkδf+fk | (5) |
| --- | --- | --- |

where δ⋅ denotes the Dirac delta function. This formulation demonstrates the mechanism of noise rejection: if the modulation frequencies {fk} are designed to avoid the dominant bands of environmental noise without mutual interference, the projection value zk is retrieved by extracting the magnitudes at the corresponding carrier frequencies:

|  | Yfk=12zkak,1⇒zk=2Yfkak,1 | (6) |
| --- | --- | --- |

To ensure strict orthogonality and avoid mutual interference within a physical optoelectronic system, the carrier allocation must satisfy three boundary conditions: spectral separability (fi-fj>Δf), harmonic interference avoidance (fj≠n⋅fi), and intermodulation avoidance (fi±fj≠fk) (Mathematical derivations of these constraints and the corresponding switch search algorithm are detailed in Supplement Sections 2 and 3).

The resilience of OFE against additive white Gaussian noise (AWGN) fundamentally stems from its multiplex advantage over conventional TDM. For a total acquisition window T comprising N discrete samples, a K-channel TDM scheme allocates a disjoint time slot τ= T/K per channel. Utilizing only N/K samples for signal averaging, its projection noise variance scales as σTDM2∝K/N. Conversely, OFE modulates all K channels continuously across T. Coherent frequency-domain demodulation acts as a matched filter integrating over all N samples, bounding the noise variance to σOFE2∝1/N. This inherent variance ratio of K mathematically dictates a reduction in demodulation error by a factor of K under identical acquisition time. This theoretical multiplexing gain—conceptually analogous to Fellgett's advantage—constitutes the bedrock of OFE's noise immunity under detector-noise-limited conditions.

The aggregated DC bias ak,0 inherent to OFE elevates the shot noise baseline, potentially offsetting Fellgett's advantage in shot-noise-limited scenarios. However, for macroscopic sensing where ambient interference and AWGN dominate, the matched filtering of OFE effectively suppresses broadband noise. This ensures the multiplexing gain remains dominant, preserving high-fidelity reconstruction even in noisy, unconstrained environments.

τ∝∝Once the projection values for all K channels are retrieved, they are aggregated to form the measurement vector z=z1,z2,…,zKT. To recover the spectrum Rλ of the target, we discretize the spectral range λ1,λ2 into N bands and model the measurement process as a system of linear equations:

|  | z=Wr+e | (7) |
| --- | --- | --- |

where r∈RN represents the ground truth (GT) spectrum, and W∈RK×N is the pre-calibrated system matrix determined by spectral sensing system. The term e accounts for the system noise and measurement residuals. Consequently, spectral reconstruction is inherently an ill-posed inverse problem. To ensure physical fidelity and numerical stability, we recast the reconstruction as a constrained convex optimization problem. We employ a projected gradient descent (PGD) algorithm to minimize a composite objective function that incorporates both physical boundary constraints and adaptive spectral priors:

|  | r=min0⪯r⪯112z-Wr22+Rrarg minrz-Wr22 | (8) |
| --- | --- | --- |

where Rr is a generalized regularization term adaptively selected based on the physical characteristics of the target spectra. For targets with distinct narrowband feature, we apply a sparsity prior to resolve fine spectral structures, defining Rr=β1r. For targets characterized by continuous, broadband spectral envelopes, we enforce smoothness and structural continuity by defining Rr=β22r22+β32Dr22, where D∈RN-1×N is the first-order difference operator. Detailed hyperparameter configurations and the specific algorithmic implementation are cataloged in Supplement Section 1.

Experiment

To evaluate the OFE framework, we construct a programmable prototype utilizing a customized LED array with K=12 discrete narrowband channels spanning 400 nm to 700 nm (spectral profiles and radiant fluxes are pre-calibrated as shown in Fig. S2).

To isolate and quantify the fundamental multiplexing advantage of the OFE framework against AWGN, we establish a theoretical benchmark. In a pure AWGN regime with a flat noise power spectral density, the specific spectral placement of orthogonal carriers mathematically yields identical variance reduction. Therefore, to ensure baseline orthogonality without introducing redundant hardware-specific combinatorial complexities, we employ a straightforward sub-octave linearly spaced carrier set for all numerical TDM-vs-OFE comparative simulations. Specifically, this frequency set (Strategy 1) spans from fmin=380 kHz to fmax=490 kHz with a uniform frequency resolution of Δf = 10kHz(detailed algorithmic implementation is provided in Supplement Section 3). Conversely, to ensure an equivalent temporal and energy budget for comparison, the TDM benchmark uniformly partitions the identical 100 μs acquisition window into K=12 sequential time slots (≈8.33 μs per channel), with each channel operating at a constant direct current (DC) baseline.

The numerical simulations are conducted utilizing the CAVE multispectral database. The reported spectral reconstruction RMSE and demodulation errors in Fig. 2 represent the statistical average across all spatial pixels from the 29 diverse hyperspectral images contained within the dataset.

Figure 2. Numerical benchmarking of the OFE and TDM frameworks. (a) Time-domain waveforms under AWGN (SNR = 25 dB). (b) Spectral reconstruction RMSE (top) and demodulation error (bottom) versus AWGN SNR. Data points represent statistical averages across the CAVE database. (c) Time-domain waveforms under multitone interference (Δfguard=10 kHz, SNR = -10 dB). The 10 kHz offset corresponds to the 1/T frequency resolution limit of the integration window. (d) RMSE versus multitone interference SNR at varied frequency offsets. TDM yields a single baseline (grey dashed line) due to the absence of frequency-domain selectivity. OFE achieves interference rejection exclusively when Δfguard meets the 10 kHz integer-cycle threshold.

Note that the DC bias (ak,0) requisite for physical LED emission is pre-filtered via AC-coupling, presenting only the zero-mean information-bearing composite signal. Figure 2(a) illustrates the time-domain waveforms of both schemes under AWGN (SNR = 25 dB). The OFE signal exhibits a continuous, dense composite waveform characteristic of parallel frequency multiplexing, whereas TDM presents discrete, sequential step-like features. Figure 2(b) quantifies the reconstruction accuracy across a broad SNR spectrum. To directly link the hardware acquisition quality to the algorithmic reconstruction, we track both the spectral reconstruction RMSE (left axis) and the demodulation error (right axis, defined as the relative Euclidean distance z-z2/z2 between the extracted and ground-truth spectral projections). Across all evaluated noise levels, OFE achieves lower errors than TDM. The observed separation between the curves rigorously aligns with the theoretical K (approximately 3.46-fold for K=12) multiplexing advantage derived in our principle section.

Beyond broadband AWGN, practical systems frequently encounter structured interference. To isolate the system's dynamic response to such artifacts, we maintain the identical TDM and OFE configurations used in the AWGN benchmark, but substitute the broadband noise model with severe multitone interference operating at an aggressive SNR of -10 dB. We parameterize the frequency offset between each interfering tone and its adjacent OFE carrier as the guard band, Δfguard=fnoise-fcarrier.

Figures 2(c) and 2(d) elucidate the interference rejection mechanism. In the time-domain visualization (Fig. 2c), multitone interference (SNR = -10 dB) at an offset of Δfguard=10 kHz induces high-frequency ripples. Because this offset precisely matches the frequency resolution limit (1/T) of the acquisition window, the noise harmonic completes integer cycles within the integration period. This satisfies the orthogonality condition, enabling its complete elimination during frequency-domain demodulation. Figure 2(d) quantifies this threshold behavior. TDM lacks frequency-domain selectivity; consequently, its time-domain integration degrades uniformly across all evaluated offsets, collapsing into a single baseline (grey dashed line). For OFE, noise immunity is governed by the 1/T threshold. Offsets below this limit (Δfguard=1 and 5 kHz) violate the integer-cycle condition, causing spectral leakage and elevated RMSE. Once the offset reaches the 10 kHz threshold, orthogonality is restored, yielding a consistently low reconstruction error independent of the interference intensity.

Having established the theoretical noise immunity, we develop a physical optoelectronic prototype to empirically validate the OFE framework. The illumination module comprises the identical array of K=12 discrete narrowband LEDs (spanning 400 nm to 700 nm) utilized in the numerical simulations. This configuration ensures that the physical system matrix W maps to our theoretical benchmark. Figure 3(a) illustrates the versatile optoelectronic prototype configured for both reflective and transmissive modalities (a photograph of the physical benchtop setup is provided in Supplement Section X, Fig. SX).

To assess the system's baseline reconstruction capabilities without external electrical interference, we configure the hardware for a high-speed snapshot acquisition (Ts=250 μs). As established by the empirical benchmark (Fig. SX in Supplement), this stringent temporal boundary dictates a coarse theoretical frequency resolution of Δf = 4 kHz. To multiplex all K=12 carriers within the available hardware bandwidth under this constraint, we deployed the broadband greedy search algorithm (Strategy 2). Unlike the dense, linearly spaced configuration used in the pure AWGN benchmark, this strategy yields a non-uniformly distributed carrier set (Hopt∈30,410 kHz) that structurally navigates the complex intermodulation landscape, maximizing spectrum utilization in a clean electromagnetic environment (the exact numerical array is detailed in Supplement Section 3).

Figures 3(b) and 3(c) show the spectral reconstruction fidelity under this snapshot regime. Employing the reflective optical path, Figure 3(b) demonstrates the recovery of continuous broadband envelopes from a standard 24-patch color checker. The system yields a mean RMSE of 0.0422, which, as visually corroborated by the insets, translates into consistent sRGB colorimetric reproduction across the diverse reflection profiles. Utilizing the transmissive path, Figure 3(c) validates the recovery of isolated 10 nm FWHM narrowband filters. The extracted spectral vectors successfully resolve these sharp peaks, preserving the physically discontinuous features.

ΔΔΔΔΔΔΔΔ

N/KK KK≈3.5K

ΔfΔfΔf

Δf

Figure 3. Figure 3. Experimental validation of the physical OFE prototype in the high-speed regime (Ts=250 μs, Strategy 2). (a) Schematic of the experimental setup for physical spectrum acquisition. (b) Reconstructed surface reflectance of the 24-patch color checker. Open circles denote the OFE prediction (Pred.) against the ground truth (GT, dashed lines), yielding an average RMSE of 0.0422. Insets display the corresponding sRGB color renderings. (c) Reconstruction profiles of isolated 10 nm FWHM narrowband spectral peaks. TS=

To evaluate the framework's survivability, we introduced a 50 kHz switched-mode power supply (SMPS) noise source. For this rigorous edge-case stress test, we utilized the transmissive optical path and selected a representative broadband profile (transmissive colorchecker patch #15) as the fixed physical probe for all subsequent evaluations in Figure 4.

To quantify the system's performance under such extreme conditions, we define an effective frequency-domain SNR. As detailed in Supplement Section X, this metric strictly compares the power of the fundamental carriers against the broad environmental noise floor. By excluding the deterministic, higher-order hardware harmonics from the noise calculation, we isolate the true stochastic impact of the SMPS interference, providing an assessment of the framework's environmental robustness. Figures 4(a) and 4(b) contrast the time-domain composite signals under pristine (SNR = 23.11 dB) and extreme interference (SNR = -11.77 dB) conditions. In Figure 4(b), the high-energy SMPS noise entirely masks the periodic modulation envelope.

≈≈However, the system's resilience is grounded in the strict frequency-domain separation depicted in Figures 4(c) and 4(d). The SMPS interference injects massive deterministic harmonics that dominate the low-frequency spectrum up to approximately 350 kHz (Fig. 4d). Because we dynamically deployed the sub-octave configuration (Strategy 1), the OFE carriers remain rigorously isolated within the higher-frequency 380–490 kHz band (Fig. 4c). This structural frequency isolation ensures that despite the severe time-domain distortion observed in Fig. 4b, the orthogonal carriers remain uncontaminated prior to the demodulation stage.

Figure 4(e) quantifies the aggregate statistical tolerance to this severe environmental noise. The mean RMSE exhibits a distinct robustness plateau (0.0261≈) despite a 30 dB degradation in the external SNR, with performance breakdown initiating only when the SNR breaches the ≈ -12 dB threshold. To visualize the underlying structural degradation trajectory at this boundary, Figure 4(f) maps the reconstructed spectra of the probe across the evaluated noise levels. Down to the extreme -11.77 dB threshold, the extracted curves strictly track the physical ground truth (solid black line, validating the framework's robustness within the sub-noise regime.Hopt

Figure 4. Experimental robustness evaluation under 50 kHz SMPS interference (Ts=1000 μs, Strategy 1). Time-domain composite signals under (a) high SNR (23.11 dB) and (b) extreme low SNR (-11.77 dB). (c) Frequency-domain spectrum of the clean OFE signal, with red triangles marking the sub-octave carriers (380–490 kHz). (d) Frequency-domain spectrum of the interference environment, illustrating the physical isolation between the low-frequency SMPS harmonics and the high-frequency OFE signal area. (e) Reconstruction RMSE versus SNR, demonstrating a robustness plateau up to the -12 dB threshold. (f) Reconstructed transmittance spectra across varying noise levels against the ground truth (solid black line).

Conclusion and discussion

In this work, we propose OFE architecture to resolve the speed-SNR trade-off inherent in conventional TDM. By shifting from serial scanning to parallelized frequency-domain encoding, the system achieves single-shot spectral reconstruction in 250 μs. Crucially, unlike standard lock-in techniques[15], OFE employs a strategic frequency allocation scheme that exploits the spectral gaps of structured environmental noise. This mechanism ensures robust retrieval (RMSE ≈ 0.0261) even under harsh interference with an SNR of -11.77 dB (at a 1000 μs sampling time), effectively immunizing the system against common artifacts like PWM dimming and SMPS ripples.

However, the performance boundaries of OFE are strictly dictated by hardware constraints rather than the encoding strategy itself. The acquisition speed T is governed by the time-frequency uncertainty principle, T ≈1/Δf. While our prototype is limited to 250 μs by the 500 kHz bandwidth of LED drivers, the framework holds immense potential for ultrafast spectroscopy. By transitioning to high-speed laser diodes modulated at hundreds of megahertz or even gigahertz, the system can theoretically scale to nanosecond-level acquisition times[16], allowing for a dramatic increase in the number of spectral channels without compromising temporal resolution.

To extend the system's capabilities, the OFE architecture offers versatile scalability across spectral, spatial, and functional dimensions. First, the spectral range is not confined to the visible band; by incorporating appropriate sources and detectors, the framework can be readily expanded to the ultraviolet or infrared regions[17], enabling molecular fingerprinting in complex environments. Second, coupling OFE with event cameras[18] (capable of microsecond-level temporal resolution) or SPAD arrays[19] (achieving picosecond to nanosecond-level precision) facilitates hyperspectral imaging. Finally, the active nature of OFE opens new possibilities for real-time adaptive sensing. Unlike passive spectrometers, the illumination patterns in OFE can be dynamically optimized based on initial measurements to maximize the reconstruction fidelity for specific spectral features[20], establishing a closed-loop framework for intelligent optical sensing.

References

- Cui X, Nigmatulin F, Wang L, et al. Miniaturized spectral sensing with a tunable optoelectronic interface[J]. Science advances, 2025, 11(4): eado6886.

- Zhang D, Hou L, Lv L, et al. Precision agriculture: temporal and spatial modeling of wheat canopy spectral characteristics[J]. Agriculture, 2025, 15(3): 326.

- Oberski T, Walendzik B, Szejnfeld M. The monitoring of macroplastic waste in selected environment with UAV and multispectral imaging[J]. Sustainability, 2025, 17(5): 1997.

- Zou X, Shen Y, Jiao C, et al. Label-free resonance Raman spectral imaging reveals magnesium microsphere therapy attenuates oxidative damage in knee osteoarthritic rats[J]. Spectrochimica Acta Part A: Molecular and Biomolecular Spectroscopy, 2026: 127743.

- Liu W, Xu S, Lee C. Ultracompact on-chip spectral shaping using pixelated nano-opto-electro-mechanical gratings[J]. Science, 2025, 389(6762): 806-810.

- Grotevent M J, Yakunin S, Bachmann D, et al. Integrated photodetectors for compact Fourier-transform waveguide spectrometers[J]. Nature Photonics, 2023, 17(1): 59-64.

- Hong D, Li C, Yokoya N, et al. Hyperspectral imaging[J]. Nature Reviews Methods Primers, 2026, 6(1): 19.

- Cui K, Huang Y. On-chip spectral imaging and sensing transition towards marketable technologies[J]. Nature Reviews Electrical Engineering, 2025, 2(3): 151-152.

- Chen Z, Qu G, He Z, et al. Single‐Pixel Multispectral Imaging With Customized Pixelated Spectral Mask Plate[J]. Laser & Photonics Reviews, 2026: e03207.

- Hong D, Zhang B, Li X, et al. SpectralGPT: Spectral remote sensing foundation model[J]. IEEE transactions on pattern analysis and machine intelligence, 2024, 46(8): 5227-5244.

- Hua X, Wang Y, Wang S, et al. Ultra-compact snapshot spectral light-field imaging[J]. Nature communications, 2022, 13(1): 2732.

- Yao Z, Liu S, Wang Y, et al. Integrated lithium niobate photonics for sub-ångström snapshot spectroscopy[J]. Nature, 2025, 646(8085): 567-575.

- Zhou P, Lv Y, Zheng Y, et al. CTIS spectral image reconstruction technology based on slit-scanning architecture[J]. Optics Express, 2024, 32(14): 25276-25291.

- Adhikari S, Gross N, Wilson K S, et al. On-chip lock-in detection for ultrafast spectroscopy of single particles[J]. The Journal of Physical Chemistry C, 2024, 128(21): 8708-8715.

- Shao Z, Huang G, Mielczarek A, et al. Robust spectral sensor for standoff biometric detection[J]. Nature Sensors, 2026: 1-8.

- Wörner H J, Wolf J P. Ultrafast spectroscopy of liquids using extreme-ultraviolet to soft-X-ray pulses[J]. Nature Reviews Chemistry, 2025, 9(3): 185-199.

- Yu H, Memon M H, Yao M, et al. A miniaturized cascaded-diode-array spectral imager[J]. Nature Photonics, 2025, 19(12): 1322-1329.

- Yu B, Liang J, Wang Z, et al. Active hyperspectral imaging using an event camera[C]//Proceedings of the Computer Vision and Pattern Recognition Conference. 2025: 929-939.

- Cao Y, Bao H, Peng L, et al. Single-Photon-Camera-Based Time and Spatially Resolved Electroluminescence Spectroscopy for Micro-LED Analysis[J]. ACS Photonics, 2024, 11(2): 731-736.

- Yuan Y, Sun R, Xie X, et al. Multispectral imaging colorimetric measurement method with illumination-adaptive white calibration[J]. JOSA A, 2025, 42(7): 898-907.